Alarm

S/R ENBL

Data Gate

Tm In

µG

Disable

µG Trig

Fast Veto

S/R

Trig

& Gate

Tm Out

Busy

Data

Gate

TDC

Gate

µ Gated

Ref

Gate

S/R

Aux1

(TTL)

Rate In

Reset

& Go

Preset

Counter

Out

Inhibit

GGL

fm

Pulser

D

A

C

µ Stop

Alarm

µ Gated

Enable

IMuSR

µSR VME GATE GENERATOR LOGIC Module

General Description

The VME Gate Generator Logic Module (GGL) was designed to generate the following three gate pulses.

##### Tm In                                                                    >Δ+δ1+δ2

Data Gate                           Δ

TDC Gate                          Δ+δ1

Ref Gate                          Δ+δ1+δ2

20ns < Δ < 20.47µs          20ns < δ1, δ2 < 1.27µs

Once the Data Gate, the TDC Gate and the Ref Gate has been triggered by Tm In pulse, another Tm In pulse occurring before the end of Ref Gate will not extend any of the three gate pulses. Also, a Tm In pulse will trigger the three gate pulses only after Δ+δ1+δ2 seconds from the last Tm In pulse.

In addition, a 32-bit down counter capable at running at 10MHz was implemented. An active signal will indicate that the count value has reached zero and no pulses will be allowed thru the counter. The count value can be reset at any time via the front panel interface or via VME.

A variable period and duty cycle square wave generator was implemented. The period ranges from 0.4 milliseconds to 13.107 seconds.

A 16-bit DAC was also implemented.  Six ranges are available: 0 to +5V, 0 to +10V, -2.5V to +2.5V, -5V to +5V, -10V to +10V and -2.5V to +7.5V.  All output ranges can drive up to 100 milliamperes.  Range selection is performed via a VME register.

VME Interface 	SLAVE – A16, D16, D8 (OE)

The GGL requires a 16-bit address space. Jumpers on the printed circuit board configure the base address selection.

Address Modifier Selection

The GGL will only respond to A16 address cycles.

Short Supervisory & short nonprivileged access - 0x2D, 0x29

Base Address Selection

Each jumper corresponds to address bits A15 – A5 on the VME address bus. Installing a jumper for each address bit will select a 0 (low) for the corresponding VME address bit.

Table 1 Base Address Selection

| Jumpers Installed (X) |  |  |  |  |  |  |  |  |  |  | Address Range |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| A15 | A14 | A13 | A12 | A11 | A10 | A9 | A8 | A7 | A6 | A5 |  |
| X | X | X | X | X | X | X | X | X | X | X | 0000 – 001F |
| - | X | X | X | X | X | X | X | X | X | X | 8000 – 801F |
| - | - | X | X | X | X | X | X | X | X | X | C000 – C01F |
| - | - | - | X | X | X | X | X | X | X | X | E000 – E01F |
| - | - | - | - | X | X | X | X | X | X | X | F000 – F01F |
| - | - | - | - | - | X | X | X | X | X | X | F800 – F81F |
| - | - | - | - | - | - | X | X | X | X | X | FC00 – FC1F |

## Input Description

Inputs are: Tm In, µG Disable, µG Trig Fast Veto, S/R Trig & Gate, Rate In, Reset & Go.

All inputs accept NIM levels. Tm In is not 50 Ω terminated. To see a true NIM level, the user must terminate at Tm Out or plug Tm Out into a 50 Ω terminated NIM input.

## Digital Output Description

Outputs are: Tm Out, Busy, µ Stop, Data Gate, TDC Gate, Alarm, µ Gated, Ref Gate, S/R Enable, Aux1, fm Pulser, Preset Counter Out, Inhibit.

All digital outputs with the exception of AUX1 drive NIM levels. AUX1 is a TTL output with a 50 Ω drive.

## µG Trig Fast Veto Setup

The minimum setup time is 3ns with respect to the Tm In signal. The pulse width of the µG Trig Fast Veto should be set to 5ns (assuming 3ns setup time).  Veto efficiency goes to zero with a 2ns setup time.

Table 2 Default Register Values

| Function | Address | Value |
| --- | --- | --- |
| ##### Main Delay (MSB) | ##### 00 | ##### 03 |
| ##### Main Delay (LSB) | ##### 01 | ##### E8 |
| ##### Delta 1 Delay | ##### 02 | ##### 1E |
| ##### Delta 2 Delay | ##### 03 | ##### 32 |
| ##### VME S/R Enable | ##### 04 | ##### 00 |
| ##### DAC Range Control | ##### 05 | ##### 00 |
| ##### DAC Setpoint (MSB) | ##### 06 | ##### 00 |
| ##### DAC Setpoint (LSB) | ##### 07 | ##### 00 |
| ##### Counter Preset Value (MSB, Byte 3) | ##### 08 | ##### 00 |
| ##### Counter Preset Value (Byte 2) | ##### 09 | ##### 98 |
| ##### Counter Preset Value (Byte 1) | ##### 0A | ##### 96 |
| ##### Counter Preset Value (LSB, Byte 0) | ##### 0B | ##### 80 |
| ##### Counter Readback Value (MSB, Byte 3) | ##### 0C | ##### 00 |
| ##### Counter Readback Value (Byte 2) | ##### 0D | ##### 98 |
| ##### Counter Readback Value (Byte 1) | ##### 0E | ##### 96 |
| ##### Counter Readback Value (LSB, Byte 0) | ##### 0F | ##### 80 |
| ##### Pulser HI Time Set Value (MSB) | ##### 10 | ##### 00 |
| ##### Pulser HI Time Set Value (LSB) | ##### 11 | ##### 02 |
| ##### Pulser LO Time Set Value (MSB) | ##### 12 | ##### 00 |
| ##### Pulser LO Time Set Value (LSB) | ##### 13 | ##### 02 |
| ##### Pulser enable | ##### 14 | ##### 00 |
| ##### VME GGL Alarm | ##### 15 | ##### 00 |
| ##### Not Used | ##### 16 | ##### 00 |
| ##### Not Used | ##### 17 | ##### 00 |
| ##### Not Used | ##### 18 | ##### 00 |
| ##### Not Used | ##### 19 | ##### 00 |
| ##### Not Used | ##### 1A | ##### 00 |
| ##### Not Used | ##### 1B | ##### 00 |
| ##### Not Used | ##### 1C | ##### 00 |
| ##### Counter Reset | ##### 1D | ##### 00 |
| ##### Not Used | ##### 1E | ##### 00 |
| ##### VME Register Reset | ##### 1F | ##### 00 |

## 11-Bit Δ Delay Register:   00, 01   (Default: 03E8H)

These registers will respond to a byte or a word access.

| ADR | $xxxxxx00 (Bits 15-8) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BIT | 15 | 14 | 13 | 12 | 11 | 10 | 9 | 8 |
| OPER | R | R | R | R | R | R/W | R/W | R/W |
| RESET | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 |

| ADR | $xxxxxx01 (Bits 7-0) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BIT | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0 |
| OPER | R/W | R/W | R/W | R/W | R/W | R/W | R/W | R/W |
| RESET | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 |

Δ delay = register value x 10ns

By default, Δ delay is 10µs. Value in register MUST be greater than or equal to 2.

## 7-Bit δ1 Delay Register:   02 (Default: 1EH)

This register will respond to a byte or a word access.

| ADR | $xxxxxx02 (Bits 7-0) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BIT | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0 |
| OPER | R | R/W | R/W | R/W | R/W | R/W | R/W | R/W |
| RESET | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 |

δ1 delay = register value x 10ns

By default, δ1 delay is 300ns. Value in register MUST be greater than or equal to 2.

## 7-Bit δ2 Delay Register:   03 (Default: 32H)

This register will respond to a byte or a word access.

| ADR | $xxxxxx03 (Bits 7-0) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BIT | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0 |
| OPER | R | R/W | R/W | R/W | R/W | R/W | R/W | R/W |
| RESET | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 |

Δ2 delay = register value x 10ns

By default, δ2 delay is 500ns. Value in register MUST be greater than or equal to 2.

## VME S/R Enable Register:   04   (Default: 00H)

This register will respond to a byte or a word access.

| ADR | $xxxxxx04 (Bits 7-0) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BIT | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0 |
| OPER | R | R | R | R | R | R/W | R/W | R/W |
| RESET | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |

AUX1 = bit1 xor (not (Ref Gate and bit0))

S/R Enable = bit0 xor bit2

## DAC Range Control Register:   05   (Default: 00H)

This register will respond to a byte or a word access.

| ADR | $xxxxxx05 (Bits 7-0) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BIT | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0 |
| OPER | R | R | R | R | R | R/W | R/W | R/W |
| RESET | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |

The DAC output range is +5V with an output voltage of 0V when powered initially.

Bit3..Bit 0:	000B:	0 to +5V.

001B: 	0 to +10V.

010B:	-5V to +5V.

011B:	-10V to +10V

100B:	-2.5V to +2.5V

101B:	-2.5V to +7.5V

## DAC Setpoint Register:   06, 07   (Default: 0000H)

These registers will respond to a byte or a word access.

| ADR | $xxxxxx06 (Bits 15-8) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BIT | 15 | 14 | 13 | 12 | 11 | 10 | 9 | 8 |
| OPER | R/W | R/W | R/W | R/W | R/W | R/W | R/W | R/W |
| RESET | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |

| ADR | $xxxxxx07 (Bits 7-0) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BIT | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0 |
| OPER | R/W | R/W | R/W | R/W | R/W | R/W | R/W | R/W |
| RESET | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |

|  | 0 to +5V | 0 to +10V | -5V to +5V | -10V to +10V | -2.5V to +2.5V | -2.5V to +7.5V |
| --- | --- | --- | --- | --- | --- | --- |
| ##### FFFFH | +5.00V | +10.0V | +5.00V | +10.0V | +2.50V | +7.50V |
| C000H | +3.75V | +7.50V | +2.50V | +5.00V | +1.25V | +5.00V |
| 8000H | +2.50V | +5.00V | 0V | 0V | 0V | +2.50V |
| 4000H | +1.25V | +2.50V | -2.50V | -5.00V | -1.25V | 0V |
| 0000H | 0V | 0V | -5.00V | -10.0V | -2.50V | -2.50V |

32-Bit Down Counter Preset Register:   08,09,0A,0B   (Default: 00989680H)

These registers will respond to a byte or a word access.

| ADR | $xxxxxx08 (Bits 31-24) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BIT | 31 | 30 | 29 | 28 | 27 | 26 | 25 | 24 |
| OPER | R/W | R/W | R/W | R/W | R/W | R/W | R/W | R/W |
| RESET | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |

| ADR | $xxxxxx09 (Bits 23-16) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BIT | 23 | 22 | 21 | 20 | 19 | 18 | 17 | 16 |
| OPER | R/W | R/W | R/W | R/W | R/W | R/W | R/W | R/W |
| RESET | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 |

| ADR | $xxxxxx0A (Bits 15-8) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BIT | 15 | 14 | 13 | 12 | 11 | 10 | 9 | 8 |
| OPER | R/W | R/W | R/W | R/W | R/W | R/W | R/W | R/W |
| RESET | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 |

| ADR | $xxxxxx0B (Bits 7-0) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BIT | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0 |
| OPER | R/W | R/W | R/W | R/W | R/W | R/W | R/W | R/W |
| RESET | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |

The default value is 10 million counts. The counter will stop when the count value has reached zero. The user can reset of the counter thru VME or thru the front panel interface.

## 32-Bit Down Counter Readback Register:   0C,0D,0E,0F   (Default: 00989680H)

These registers will respond to a byte or a word access.

| ADR | $xxxxxx0C (Bits 31-24) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BIT | 31 | 30 | 29 | 28 | 27 | 26 | 25 | 24 |
| OPER | R | R | R | R | R | R | R | R |
| RESET | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |

| ADR | $xxxxxx0D (Bits 23-16) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BIT | 23 | 22 | 21 | 20 | 19 | 18 | 17 | 16 |
| OPER | R | R | R | R | R | R | R | R |
| RESET | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 |

| ADR | $xxxxxx0E (Bits 15-8) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BIT | 15 | 14 | 13 | 12 | 11 | 10 | 9 | 8 |
| OPER | R | R | R | R | R | R | R | R |
| RESET | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 |

| ADR | $xxxxxx0F (Bits 7-0) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BIT | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0 |
| OPER | R | R | R | R | R | R | R | R |
| RESET | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |

These register displays the current count value of the 32-bit down counter.

## Pulser HI Time Setpoint Register:   10, 11   (Default: 0002H)

These registers will respond to a byte or a word access.

| ADR | $xxxxxx10 (Bits 15-8) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BIT | 15 | 14 | 13 | 12 | 11 | 10 | 9 | 8 |
| OPER | R/W | R/W | R/W | R/W | R/W | R/W | R/W | R/W |
| RESET | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |

| ADR | $xxxxxx11 (Bits 7-0) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BIT | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0 |
| OPER | R/W | R/W | R/W | R/W | R/W | R/W | R/W | R/W |
| RESET | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 |

By default, HI time is 200µs. Value in registers MUST be greater than or equal to 2.

## Pulser LO Time Setpoint Register:   12, 13   (Default: 0002H)

These registers will respond to a byte or a word access.

| ADR | $xxxxxx12 (Bits 15-8) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BIT | 15 | 14 | 13 | 12 | 11 | 10 | 9 | 8 |
| OPER | R/W | R/W | R/W | R/W | R/W | R/W | R/W | R/W |
| RESET | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |

| ADR | $xxxxxx13 (Bits 7-0) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BIT | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0 |
| OPER | R/W | R/W | R/W | R/W | R/W | R/W | R/W | R/W |
| RESET | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 |

By default, LO time is 200µs. Value in registers MUST be greater than or equal to 2.

## Pulser Enable Register:   14   (Default: 00H)

This register will respond to a byte or a word access.

| ADR | $xxxxxx14 (Bits 7-0) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BIT | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0 |
| OPER | R | R | R | R | R | R | R | R/W |
| RESET | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |

The Pulser is disabled when powered initially.

Bit 0:	0:	Disabled.

1: 	Enabled.

## Period = ( HI Time Setpoint + LO Time Setpoint ) x 0.0001 seconds.

## Alarm Register:   15  (Default: 00H)

This register will respond to a byte or a word access.

| ADR | $xxxxxx1D (Bits 7-0) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BIT | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0 |
| OPER | R | R | R | R | R | R | R | R/W |
| RESET | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |

Bit 0:	0:	Disabled.

1: 	Enabled.

## 32-Bit Down Counter Reset Register:   1D  (Default: 00H)

This register will respond to a byte or a word access.

| ADR | $xxxxxx1D (Bits 7-0) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BIT | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0 |
| OPER | R/W | R/W | R/W | R/W | R/W | R/W | R/W | R/W |
| RESET | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |

Write cycle will reload 32-bit down counter with the value in the counter set registers. Value read from this register is always zero.

## VME Reset Control Register:   1F  (Default: 00H)

This register will respond to a byte or a word access.

| ADR | $xxxxxx1F (Bits 7-0) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BIT | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0 |
| OPER | R/W | R/W | R/W | R/W | R/W | R/W | R/W | R/W |
| RESET | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |

Write cycle will reset all VME registers to default values. However, the DAC output voltage and range will remain unchanged and not reflected in the register values.

## Block Diagram of Module Operation